[UNIVERSITY LOGO]

[DEPARTMENT]
[INSTITUTION]

TRANSCRANIAL MAGNETIC STIMULATION (TMS)
Rules of Operation

[VERSION HISTORY]
[LATEST DATE]

This document has been reviewed and approved by the [ETHICS REFERENCE]
# Contents

1. Introduction
  1. The TMS committee
  2. Internal approval of TMS studies
2. TMS technique
2.1. Outline of technique
2.2. Types of TMS
2.3. Uses of TMS
Safety Guidelines
3.1. Participant screening
3.2. Potential side effects
3.3. Precautions
3.4. Authorised Users
3.5. Managing emergencies

Appendices
Participant screening form
Rules for the administration of screening forms
Drugs which are potential hazards for TMS
Participant information sheet
Participant consent form
Follow-up form (to be completed after every TMS session)
TMS users test
Introduction
This document outlines guidelines for the use of transcranial magnetic stimulation (TMS) for research purposes at the [DEPARTMENT; INSTITUTION].

Any person using TMS must read and adhere to these guidelines.

This document was compiled from currently available literature regarding safety and ethical guidelines for the use of TMS; the main reference source being:
Rossi S, Hallett M, Rossini PM, Pascual-Leone A, & The Safety of TMS Consensus Group (2009). Safety, ethical considerations, and application guidelines for the use of transcranial magnetic stimulation in clinical practice and research. Clinical Neurophysiology, 120(12):2008-2039.

A comment has been published:
# Nyffeler T, Müri R (2010). Comment on: Safety, ethical considerations, and application guidelines for the use of transcranial magnetic stimulation in clinical practice and research, by Rossi et al. (2009). Clinical Neurophysiology, 121(6):980.

An update to the suggested standard screening questionnaire has also recently been published:
Rossi S, Hallett M, Rossini PM, Pascual-Leone A (2011) Screening questionnaire before tms: an update. 122(8):1686.

Other more recent relevant publications include:
Tringali S, Perrot X, Collet L, Moulin A (2012). Repetitive transcranial magnetic stimulation: hearing safety considerations. Brain Stimulation, 5(3):354-363.

In addition to these rules, all TMS research projects must be approved by the [ETHICS COMMITTEE].

  1. The TMS committee

The TMS Committee will consist of the following permanent members of academic staff:
- [COMMITTEE MEMBER 1]
- [COMMITTEE MEMBER 2]
- [COMMITTEE MEMBER 3]
- [COMMITTEE MEMBER 4]

In addition, one member of the committee will be drawn from the academic staff at the [DEPARTMENT]. This additional member is required not to be a user of TMS nor to be collaborating on any TMS-related projects with members of the committee. This person may be drawn from the academic staff who sit on the [ETHICS COMMITTEE].

1.2 Internal approval of TMS studies
Researchers wishing to do TMS studies in [DEPARTMENT] are required to submit a Project Proposal Form online, and to give a [SEMINAR] prior to the start of their study. The TMS Committee will monitor all TMS studies in [DEPARTMENT].
1. TMS technique
  1. Outline of technique
TMS is a non-invasive method used to stimulate the human brain and nervous system for the purposes of scientific and medical research and for therapeutic interventions. A magnetic field is generated when an electric current flows through a wire. A rapid variation of an electrical current in a coil of wires has the potential to induce electric current in a nearby coil of wires. This principle, of electromagnetic induction, originally observed by Faraday in 1831 (Barker, 1999), is the basis of the TMS technique. The electric current flow in the TMS coil generates a magnetic field that passes across the scalp and skull, and induces a flow of electric current in the underlying brain tissue. This electric current accumulates in the membranes of the neurons (nerve cells) underneath the coil, and can result in these neurons firing an electrical impulse. Because the brain or nervous system is stimulated indirectly via magnetic induction, rather than directly by electrical current, the TMS technique is more focal and much less uncomfortable than previous brain stimulation methods such as transcranial electrical stimulation (TES).

The TMS equipment consists of two main parts: a charging system with one or more capacitors that produce a rapid discharge of a high amplitude electric current, and a stimulating coil (often figure-of-eight shaped) that delivers the magnetic pulses to the brain region of interest.

Throughout this document, TMS will be used to refer not only to stimulation of the brain, but also to stimulation of the brain stem, spinal cord, peripheral nerves and muscles. The ethical and safety implications apply equally to all these types of nervous tissue stimulation. For simplicity, the word 'brain' or 'cortex' will be used in place of all nervous tissue.

  1. Types of TMS
Four main stimulation parameters define the different types of TMS currently in use.

2.2.1. Intensity: The intensity of the electrical current, and the resulting strength of the magnetic field, are typically expressed as a percentage of the maximum output of the stimulation equipment. For most equipment, 100% corresponds to a magnetic field strength of 1.5 to 2.2 Tesla. As an alternative, stimulation intensity is often expressed relative to a measurable physiological baseline in each individual participant. Such baselines could include, but are not limited to: a) The resting or active motor thresholds (RMT or AMT) for eliciting muscle twitches following TMS over the motor cortex; b) The phosphene threshold (PT) for eliciting visual phosphenes (flashes of light) following TMS over the visual cortex. The optimal physiological baseline, if used, should be determined for each study.

2.2.2. Frequency: When more than one TMS pulse is given at regular time-intervals, the frequency of stimulation is important. Frequencies of 0.1 or 0.2Hz (e.g., 1 pulse every five or ten seconds) or less are thought to be equivalent to single-pulse TMS, since cumulative effects of such stimulation have not been found. Frequencies between 0.2 and 1Hz are known as 'slow' or 'low frequency' repetitive TMS (rTMS), while frequencies over 1Hz are known as 'rapid' or 'high frequency' rTMS. The maximum stimulation frequency is limited by the TMS equipment, as well as by the safety guidelines.

2.2.3. Duration: In rTMS, the duration defines the time for which a particular frequency of stimulation is maintained without a break (a 'train' of stimulation). For low frequency rTMS, this train could be up to 30 minutes or longer, while for high frequency rTMS, durations as short as fractions of a second may be used.

2.2.4. Inter-train interval: In high frequency rTMS, this is the pause between the end of one train of stimulation and the start of the next. This duration may range from fractions of a second to minutes. Forms of stimulation with short inter-train intervals constitute 'patterned rTMS'.

Certain combinations of these parameters result in the most commonly-used forms of TMS:
1. Single-pulse TMS (single pulses separated by more than 5s).
2. Paired-pulse TMS (two pulses separated by very short intervals (<1s).
3. Low frequency rTMS (a series of three or more pulses at 0.2-1Hz).
4. High frequency rTMS (a series of three or more pulses at >1Hz).
5. Patterned rTMS (e.g., 'theta-burst stimulation', TBS: repeated trains of three pulses at 50Hz, separated by an inter-train interval of several hundred ms).

Single pulse TMS has been widely used for mapping the motor and visual cortices (Hallett, 2007), and for determining the precise timing of cerebral processes. Paired-pulse (or triple-, or quadruple-pulse) TMS can be applied to a single cerebral cortical area using one coil (to measure intracortical facilitation or inhibition), or to two or more cortical areas using two or more different coils (to measure cortico-cortical interactions). rTMS can be used to stimulate brain areas over a longer period of time, typically ranging from half a second to half an hour, potentially allowing for a greater modulation of neural activity which can outlast the period of stimulation. Although, generally speaking, inhibition of a brain region is achieved by delivering low frequency stimulation and excitation by high frequency stimulation, the effect of TMS on any given region depends on many interacting factors.

In order to control for the non-specific effects of TMS, studies often include a control or placebo condition called “sham TMS”, in which fake TMS pulses are delivered, with the accompanying acoustic stimuli. Alternative control conditions may include, but are not limited to: a) Applying real TMS pulses to a different brain area thought to play no role or a different role in the processes being studied; b) Reversing the direction of current flow within the coil, which is thought to stimulate a different set of neurons under the coil; c) Discharging the TMS coil near to or touching the participant's body, but positioned in such a way as to be ineffective in stimulating neural tissue.

  1. Uses of TMS
Three general classes of TMS research study have been recognized (see Rossi et al., 2009, pp. 2028-2029, directly quoted below).

“Class 1 (direct benefit, potential high risk): studies in patients with diagnostic or therapeutic primary objective, including the development of new therapeutic indication or protocols, with potential direct individual clinical benefit. Normal subjects should not ordinarily participate in such studies, and the risk level for patients can be theoretically high for stimulation protocols that have been not yet tested for safety.

Class 2 (indirect benefit, moderate risk): studies in patients where the potential clinical benefit is more speculative or where no clinical benefit is expected, but the study is anticipated to yield valuable data for the development of treatments, safety assessment of a cortical stimulation protocol, or improved understanding of pathophysiological mechanisms of neurological or psychiatric diseases. Normal subjects may participate as control subjects. In these studies, regimens that will place subjects at significant risk of seizures or other serious adverse effects should employ only patients and not normal subjects, because exposure to adverse effects is unacceptable for normal subjects when clinical benefit is questionable.

Class 3 (indirect benefit, low risk): studies in normal subjects and patients that are expected to yield important data on brain physiology or on safety, but have no immediate relevance to clinical problems. Normal volunteers should be permitted to participate in rTMS research when it is likely to produce data that are of outstanding scientific or clinical value.”

Class 3 studies, when following the safety guidelines, can be carried out by trained professionals (e.g., including research scientists, psychologists, and technicians), without the need for a licensed physician to be present.

The majority of studies carried out in the [DEPARTMENT] will be Class 3 studies. Some Class 1 & 2 studies may be carried out, for example, as part of future collaborations between [DEPARTMENT] and medical researchers.

With regards to the definition of ‘normal’ participants, each future TMS study will have to justify their selection and screening of potential participants to the appropriate ethics committee. In some cases, this screening may require contacting the participant’s GP.
1. Safety guidelines
  1. Participant screening
TMS is considered a safe technique, assuming that safety guidelines are strictly adhered to. There are a few contraindications for TMS which may increase the risk of participants experiencing side effects. Any individual wishing to participate in a TMS experiment must be screened prior to acceptance into the study, in order to eliminate contraindications and minimise risk. The participant screening form (Appendix) must be administered and interpreted by an authorised member of staff (see Appendix, Rules for the administration of screening forms).

Wherever possible, participants for TMS studies should first undergo an MRI scan. While recognising that such an MRI scan could provide no diagnostic or definitive contra-indication information regarding TMS, researchers may find it useful in screening participants for their TMS studies, as well as for more accurate positioning of the TMS coil (neuronavigation) during the study.

  1. Potential side effects
The most common side effects of TMS are transient changes in hearing, and transient localised pain or discomfort: headache, neck pain, toothache, and paraesthesia (tingling, numbness, ‘pins and needles’) have been reported in patients. Transient hearing changes can occur due to the noise made by the TMS coil, which may exceed the recommended safety levels of auditory stimulation. Although these are reportedly the most common side effects, these symptoms only occur in a small proportion of individuals, and single-pulse TMS is particularly unlikely to cause adverse effects.

TMS, and particularly rTMS, has the potential to cause more serious side effects, such as seizure and syncope (fainting), however the risk of these occurring is very low, provided that safety precautions and participant screening guidelines are strictly adhered to. No such incidents have been reported in healthy individuals when following the established TMS safety guidelines. Only 4 such instances have been reported, all of which were in patients undergoing treatment, and three of whom were taking pro-epileptogenic medication (Rossi et al., 2009, p. 2019).

  1. Precautions
The following precautions must be taken for any study using TMS:

Earplugs or another form of ear protection must be worn by participants for the duration of TMS. The ear protection used must have a noise reduction rating (NRR) of at least 35. Researchers must ensure, by consulting the TMS hardware documentation, that the noise levels associated with their protocol do not exceed the recommended safety levels, as specified by the Occupational Safety and Health Authority (OSHA; guidelines at http://www.osha.gov).

High frequency rTMS protocols must not exceed the conventional parameters of stimulation (intensity, frequency, duration, or inter-train interval) reported by Rossi et al. (2009, Tables 3, 4, 5, and 6, on pages 2022, 2023, & 2029).

Visual monitoring of participants is mandatory in all TMS studies. This must include monitoring for muscle twitches related to the TMS.

During TMS, signs of the spread of cortical excitability should be considered as the earliest indication of a TMS-generated epileptogenic effect. If such signs are detected, the experiment should be stopped immediately and the participant closely monitored.

When using high frequency or patterned rTMS, additional monitoring of the spread of cortical excitability should be carried out using EMG or EEG, wherever this is feasible.

When using novel TMS protocols or not yet explored paradigms, for TMS studies using simulation parameters exceeding current safety guidelines, and for patients with additional risk, EMG or EEG monitoring is mandatory.

Credit cards, floppy disks, memory sticks, cathode ray tubes and other similarly sensitive electronic or magnetic objects should be kept at least 50cm away from the stimulating coil.

  1. Authorised Users
These are essential staff who are conversant with, and are able to put into practice, all the rules and emergency practices outlined in this document and its Appendices. They are responsible for the screening of participants to ensure that it is safe for them to take part in TMS studies, and they must be present and must be in direct supervision when all non-authorised TMS users (e.g., trainees) are performing TMS.

Authorised User status can only be granted by the TMS committee. For individuals who have not used TMS previously, or who are not deemed competent (at the discretion of the committee), authorised status can only be granted after a person has:

Thoroughly read the most recently published safety guidelines for TMS use (currently ‘Safety, ethical considerations, and application guidelines for the use of transcranial magnetic stimulation in clinical practice and research’, Rossi et al., 2009, and the updated screening guidelines, Rossi et al., 2011, as well as any future revisions or relevant commentaries).

Thoroughly read the latest version of the TMS Rules of Operation.

Thoroughly read the relevant Health and Safety guidelines and the TMS risk assessment.

Completed a written test (administered by the TMS committee, see Appendices) which covers the safety guidelines and procedures outlined in the Rules of Operation and the latest published safety guidelines.

- Completed ten experimental sessions (totalling at least 10 hours) of TMS administration under the supervision of an Authorised User, until the TMS committee is satisfied that the individual is fully competent in all relevant aspects of TMS use and safety. This training must include:
- Measurement of resting or active motor threshold using a standardised method (e.g., Awiszus, 2003; Rossi et al., 1994)
- Recording and interpretation of electromyographic (EMG) data from a hand or other muscle
- Neuronavigation using EEG 10:20 system and/or MRI scans

Completed a basic first aid course, such as Emergency First Aid At Work.

Authorised Users must be trained in the use of any new equipment, software or procedures that may be introduced from time to time.

Trained service personnel or representatives of the equipment manufacturers can operate the equipment for quality control testing, servicing and demonstration purposes. Such people may be admitted access to the equipment by Authorised Users on production of identification (and normally by arrangement with the TMS committee).

  1. Managing emergencies
First aid equipment must be available in the same room during any TMS experiment

In case of seizure or other serious side effects, emergency medical assistance must be summoned immediately (by telephoning 999)

In case of seizure or other serious side effects, follow up management of the psychosocial consequences must be provided

All adverse effects will be recorded and reported to the TMS committee, including careful documentation of machine type, coils used, TMS intensity as a function of maximum stimulator output, and total number and duration of stimulation (this documentation will be available as a matter of course for all TMS experiments, see Appendices).

Appropriate consideration of participant confidentiality must be made in the reporting of TMS-induced seizures, so as to minimise the possibility of the participant suffering negative consequences due to the denial of medical insurance or employment problems:

“It is readily imaginable that the report of a seizure in the medical record of a normal volunteer or certain patients could be misinterpreted or deliberately used as a pretext for the denial of employment or medical insurance. Subjects of research studies must be informed of this possibility, and investigators must make certain that documentation of seizures is done in such a way that jeopardizes subjects to the minimum extent possible. Additional documentary support of a healthy subject’s claim that a provoked seizure carries no adverse prognosis must be provided when appropriate.” (Rossi et al., 2009, p. 2032).